# TEMPO EC Method for automated enumeration of Escherichia coli in food – AOAC 2009.02

## SCOPE

This method is applicable for enumeration of E. coli in meat and meat products and carcass swabs.

## PRINCIPLES

TEMPO EC (Escherichia coli) is an automated test for the enumeration of E. coli in 24 hours in food products. The TEMPO EC test consists of a vial of culture medium and a ‘card’, which are specific to this test.  E coli if present reduce the substrate in the culture medium resulting in a fluorescent signal, which is detected by the TEMPO Reader. The TEMPO system estimates the number of E coli present in the original sample according to a calculation based on the Most Probable Number (MPN) method.

The enumeration of E coli is broken down into stages as follows:

### Sample preparation

Samples are diluted 1:10 in Butterfield’s buffered peptone water and are homogenize in the TEMPO bag. For carcass swabs samples can be taken directly from the swab without further dilution. TEMPO Automated System is used to reconstitute culture media (using sterile distilled water). Media and sample mixture (1 mL filtered sample in 3 mL media) are transferred into the TEMPO card (consult the manufacturer’s instructions).

### Incubation

TEMPO card with media is incubated at 35 ± 1ᵒC for 22 – 27 hours.

### Interpretation

Reading is performed by the TEMPO system by counting the number of positive wells obtained in relation to the volume of the wells and the dilution of the sample. Results are expressed in CFU per g, mL or cm2 of original sample, using MPN tables.

## CHECKLIST

| Sample Preparation | Is Butterfield’s buffered peptone water used for preparation of samples and dilutions? |  |
| --- | --- | --- |
|  | Is sterile distilled water used for reconstitution of the TEMPO media? |  |
|  | Is TEMPO media stored at 2–25ᵒC in a dark place? |  |
|  | Is the dilution used capable of detecting E. coli at the level required? |  |
|  | Is a positive control run with each batch of samples analysed (or daily)? |  |
| Incubation | Are TEMPO cards incubated for 22-27 h at 35 ± 1ᵒC? |  |
| Interpretation | Are the manufacturer’s instructions available and are they reproduced in the laboratory manual? |  |
|  | Are technicians familiar with the operation of TEMPO Automated System? |  |
|  | Is the count correctly reported ie CFU/cm2 for swab and surface samples and CFU/g for meat samples? |  |