# Approved Arrangements

For 5.3—Biosecurity containment level 3(BC3)

Requirements— Version 3.1



© Commonwealth of Australia

**Ownership of intellectual property rights**

Unless otherwise noted, copyright (and any other intellectual property rights, if any) in this publication is owned by the Commonwealth of Australia (referred to as the Commonwealth).

**Creative Commons Licence**

All material in this publication is licensed under a Creative Commons Attribution 3.0 Australia Licence, save for content supplied by third parties, logos and the Commonwealth Coat of Arms.

Creative Commons Attribution 3.0 Australia Licence is a standard form licence agreement that allows you to copy, distribute, transmit and adapt this publication provided you attribute the work. A summary of the licence terms is available from creativecommons.org/licenses/by/3.0/au/deed.en. The full licence terms are available from creativecommons.org/licenses/by/3.0/au/legalcode.

This publication (and any material sourced from it) should be attributed as: Approved Arrangements section, 2016, AA <name> - Requirements, (AA for <name> - Requirements prepared for the Department of Agriculture and Water Resources), Canberra.

**Department of Agriculture and Water Resources**

Postal address GPO Box 858

Canberra ACT 2601

Switchboard +61 2 6272 3933

Web [agriculture.gov.au](http://www.agriculture.gov.au/)

Inquiries regarding the licence and any use of this document should be sent to: [copyright@agriculture.gov.au](mailto:copyright@agriculture.gov.au).

The Australian Government acting through the Department of Agriculture and Water Resources has exercised due care and skill in the preparation and compilation of the information and data in this publication. Notwithstanding, the department, its employees and advisers disclaim all liability, including liability for negligence, for any loss, damage, injury, expense or cost incurred by any person as a result of accessing, using or relying upon any of the information or data in this publication to the maximum extent permitted by law.

**Version control**

Updates to this document will occur automatically on the department’s website and the revision table below will list the amendments as they are approved.

| Date | Version | Amendments | Approved by |
| --- | --- | --- | --- |
| 05 Nov 2008 | 1.0 | Revised document | Co-regulation and Support Program |
| 30 Jun 2013 | 1.1 | Updated to reflect DAFF branding | Industry Arrangements Reform Program |
| 30 Mar 2015 | 1.2 | Minor update in response to restructure | Approved Arrangements section |
| 16 Jun 2016 | 2.0 | Updated to new template, references to the department and the Biosecurity Act 2015 | Approved Arrangements section |
| 24 Jun 2016 | 3.0 | Added Accredited Persons requirements | Approved Arrangements section |
| 4 July 2016 | 3.1 | Updated terminology and added another Accredited Persons requirement. | Approved Arrangements section |

Contents

[Guide to using this document 6](#_Toc485902877)

[Definitions 6](#_Toc485902878)

[Other documents 6](#_Toc485902879)

[AA Requirements 7](#_Toc485902880)

[1. Purpose 7](#_Toc485902881)

[2. Scope 7](#_Toc485902882)

[3. Requirements for approval 8](#_Toc485902883)

[4. Requirements to maintain approval 8](#_Toc485902884)

[General requirements 10](#_Toc485902885)

[5. Hygiene and isolation 10](#_Toc485902886)

[6. Biosecurity area 11](#_Toc485902887)

[7. Security 12](#_Toc485902888)

[8. Operating procedures 12](#_Toc485902889)

[9. Administration and management 13](#_Toc485902890)

[Specific requirements 16](#_Toc485902891)

[Microbiological containment—level 3 (BC3) facilities 16](#_Toc485902892)

[10. Hygiene and isolation 16](#_Toc485902893)

[11. Waste disposal 16](#_Toc485902894)

[12. Security 17](#_Toc485902895)

[13. Operational procedures 17](#_Toc485902896)

[14. Administration and management 20](#_Toc485902897)

[Indoor animal containment—level 3 (BC3) facilities 21](#_Toc485902898)

[15. Isolation and hygiene 21](#_Toc485902899)

[16. Waste disposal 21](#_Toc485902900)

[17. Security 22](#_Toc485902901)

[18. Operational procedures 23](#_Toc485902902)

[19. Administration and management 25](#_Toc485902903)

[Plant laboratory containment—level 3 (BC3) facilities 26](#_Toc485902904)

[20. Hygiene and isolation 26](#_Toc485902905)

[21. Waste disposal 26](#_Toc485902906)

[22. Security 27](#_Toc485902907)

[23. Operational procedures 27](#_Toc485902908)

[24. Administration and management 29](#_Toc485902909)

[Applicable Australian/New Zealand Standards 31](#_Toc485902910)

[Specific standards 31](#_Toc485902911)

[25. Microbiological containment—level 3 (BC3) facilities 31](#_Toc485902912)

[26. Indoor animal containment—level 3 (BC3) facilities 32](#_Toc485902913)

[27. Plant laboratory containment—level 3 (BC3) facilities 33](#_Toc485902914)

[Australian/New Zealand Standard—safety in laboratories: microbiological aspects and containment facilities (AS/NZS 2243.3:2002) 34](#_Toc485902915)

[28. Microbiological containment—level 3 (BC3) facilities 34](#_Toc485902916)

[29. Indoor animal containment—level 3 (BC3) facilities 35](#_Toc485902917)

[30. Plant laboratory containment—level 3 (BC3) facilities 36](#_Toc485902918)

[Additional requirements 38](#_Toc485902919)

[31. Specific requirements for BC3 microbiological, animal and plant facilities 38](#_Toc485902920)

[32. Specific requirements for BC3 animal and plant facilities 40](#_Toc485902921)

## Guide to using this document

This document sets out the requirements that must be met before the relevant director will consider approval for the provision of biosecurity activities under section 406 of the Biosecurity Act 2015, otherwise known as an approved arrangement (AA).

This document specifies the requirements to be met for the approval, operation and audit of this class of AA. Compliance with the requirements will be assessed by audit.

In the event of any inconsistency between these requirements and any Import Permit condition, the Import Permit condition applies. If the applicant chooses to use automatic language translation services in connection with this document, it is done so at the applicant’s risk.

Unless specified otherwise, any references to ‘the department’ or ‘departmental’ means the Department of Agriculture and Water Resources. Any references to contacting the department mean contacting your closest regional office.

Further information on AAs, regional contact details and copies of relevant AA documentation is available on the [department’s website](http://www.agriculture.gov.au/).

### Definitions

Definitions that are not contained within the Approved Arrangements Glossary can be found in the Biosecurity Act 2015 or the most recent edition of the Macquarie Dictionary.

### Other documents

The AA General Policies should be read in conjunction with these requirements. They will assist in understanding and complying with the obligations and requirements for the establishment and operation of an AA.

## AA Requirements

### Purpose

* 1. This document sets out the requirements that will achieve the structural and procedural requirements of a class 5.3 under section 406 of the Biosecurity Act 2015 (the Act).
  2. As a condition of import, the department may impose post-entry biosecurity conditions that require certain products to be restricted for use within biosecurity facilities. The purpose of approval is to satisfy the department that the facility protects Australia’s animal, plant and human health status. This is also to ensure that post-entry biosecurity procedures are followed.

### Scope

* 1. Class 5.3 AA sites utilised for goods subject to biosecurity control which pose significant risks to animals, plants or humans if pest or disease associated with them spread outside the AA site and from which significant economic impact would result in the community or environment.
  2. Class 5.3 AA sites are not approved for the distinctive needs of other biosecurity operations, except where the establishment has separate approval under another class. For example, an AA class 5.3 site is not automatically approved as a commercial fumigation facility. This would require separate class approval under class 4.6.

Note: A Biosecurity Industry Participant (BIP) may keep more than one kind of goods in the one facility, provided the applicable requirements for those kinds of goods are met.

* 1. The facility must meet the PC3 design and construction requirements as specified in Australian/New Zealand Standard TM 2243.3:2002 and 2982.1:1997. The applicable parts of these standards are outlined in Part 5 of this document.
  2. BC3 or Physical Containment (PC) Level 3 is the whole of the space approved by the department in accordance with the department’s class requirements for AA class 5.3.

A BC3 facility may incorporate access and supporting rooms and interconnecting corridors or common space areas after entering through an airlock. It may comprise a number of like rooms such as three interconnecting microbiological laboratories but does not include combinations of different types of facilities such as animal, plant or insectary facilities within a physical containment barrier.

These facilities must be physically separate from offices used by containment facility personnel.

Body showers, toilet cubicles and drinking water appliances may be included

* 1. This kind of facility is appropriate for work with imported:
* micro-organisms
* approved plant material infected with pathogens subject to biosecurity control for in vitro and in vivo use
* infected fresh or frozen fruit and vegetable samples for in vitro use
* biological material for in vivo work in animals.

This requirement is intended to apply to a wide range of different containment facilities. It is recognised that certain structural requirements, conditions and procedures apply to facilities with different functions. As such, approval as a type of AA class 5.3 will meet the requirements of class 5.1 (excluding outdoor animal facilities) and AA class 5.2 of the same type. For example, a class 5.3 Microbiological facility will also meet the requirements of class 5.1 and 5.2 Microbiological facilities.

### Requirements for approval

The applicant must provide the department with documentary evidence (certification) that the facility complies with:

* 1. The applicable design and construction standards of the Australian/New Zealand Standard TM (AS/NZS) 2982.1:1997 and 2243.3:2002 (the relevant sections of these standards are listed below in this document).

Note: The minimum requirement for obtaining this evidence is to contract a department-approved ‘third party’ assessor. A [list of department-approved ‘third party’ assessors](http://www.agriculture.gov.au/import/arrival/arrangements/assessors) can be found on the department’s website.

* 1. An air leakage rate, at a differential pressure of 200 Pa, of no more than 120L/min (upon facility commissioning).
  2. BIPs must submit a transport plan, detailing how the consignment will be taken from the port of arrival to the AA site. When developing the plan BIPs will need to ensure the following requirements are met:
* the transport route is the most direct route between the two sites
* the route taken is on sealed roads only.

### Requirements to maintain approval

* 1. Any changes to the AA site should be carried out in a manner consistent with:
* the third party certification
* the AA requirements
* the conditions of approval
* any subsequent amendments or revisions to AS/NZS 2982.1:1997 and 2243.3:2002.
  1. A change that significantly affects the overall containment system requires re-certification: this would include structural changes to 40% of the building. If a BIP has any doubt as to whether proposed changes to the physical structure of the AA site has any potential to reduce the level of biosecurity integrity, the department’s approval must be obtained before the change is implemented
  2. To ensure conformance to the AA requirements, the department must be notified in writing within 15 working days of any alterations to AA site management arrangements.
  3. Additionally, a biosecurity officer may request that documented evidence be provided for compliance with AS/NZS 2982.1:1997 and 2243.3:2002 when additions or modifications have been made to the facility.
  4. Where any structural alterations have been made to the AA site, the BIP must, with the annual approval form, provide a written declaration outlining details of the alterations made.

Note: this requires the BIP to provide the department with a contingency plan detailing how the facility will contain the biosecurity risk during alterations.

The plan may include a decontamination aspect or instructions about how the material subject to biosecurity control will be relocated into another room or facility.

* 1. At times after approval an air leakage rate of no more than 1200L/min should be maintained.
  2. To ensure that the air leakage rate is no more that 1200L/min, air leakage testing must be undertaken every three years.

## General requirements

This part outlines the general requirements that must be complied with by the BIP of any approval for a facility, irrespective of the type of facility and the containment level to which the facility is approved.

### Hygiene and isolation

* 1. Biosecurity areas must be separate from other operations within the AA sites. This can be achieved by department-approved methods. Examples of how biosecurity area separation can be defined in a particular AA class include isolation from main thoroughfares with yellow lines, structural separation, a lockable room or building, a person-proof security fence, separate benches or similar structures.

Examples of how storage separation can be achieved in a particular AA class include cupboards, cool rooms, refrigerators and freezers.

Note: Methods listed above are not applicable to all AA classes. The use of a method must be approved by the department.

* 1. For an AA class 5 to achieve the necessary separation of work and goods, it may be necessary to have cool rooms, refrigerators, freezers or other storage units located outside the area where the work is undertaken. Where this is necessary, the AA site will need to have more than one biosecurity area.
  2. For BC1 AA sites this additional biosecurity storage area may be located outside the designated facility but must be within the one physical site. To be within one physical site the facility must be within the same common boundary as the approved storage area and must be approved under the one organisation or company.
  3. Where material subject to biosecurity control is stored outside the designated facility a transfer procedure (as per 8.4) must be in place to ensure the safe movement of goods subject to biosecurity control.

Biosecurity storage areas which are located outside the building that houses the facility must be a fully enclosable space contained within walls, doors, windows, floors and ceilings. Doors and windows must be lockable and secure.

* 1. For BC2 AA sites the additional biosecurity area must be located within the building that houses the facility and where practical must be lockable. Movement procedures must be applied as per 8.4.
  2. In a class 5 AA site, where a biosecurity area is outside or separate to the area where the work is undertaken, the type of biosecurity area (such as a refrigerator or freezer) must be stated on the scale drawing.

The separation of work and goods (for example, in separate outside storage areas) is not applicable to BC3 (BC3 facilities may only have the autoclave outside the immediate facility but within the building) or BC4 facilities which must operate as a closed entity.

* 1. The AA site must be managed to ensure that effective separation is maintained between cleared imported goods, domestic goods, imported goods awaiting release from biosecurity control, and,in the case of department-approved dual import and export AA sites, export goods. BIP must also recognise that specific Import Permit conditions and inspection procedures for some commodities may apply in addition to these requirements.
  2. Effective separation of goods can be achieved by using:
* an impervious physical barrier
* other department-approved methods.
  1. Effective separation will depend on the class of goods. The methods listed are not applicable to all AA classes. Examples of effective separation for some AA classes include, but are not limited to:
* sealed containers
* storage in separate rooms
* plywood, sheet metal or heavy gauge plastic sheeting that provides complete and unbroken physical separation between consignments
* double plastic wrap including a space separation between consignments of 1.2 metres
* remaining consolidated within the shipping container.
  1. The use of a method must be approved by the department and, should cross-contamination occur, the goods shall be treated as goods subject to biosecurity control.
  2. The AA site must be managed in a way that ensures that the buildings and/or structures are maintained in a state of good repair.
  3. An effective pest control system must be in place to ensure that facilities are managed in a way that effectively isolates goods subject to biosecurity control from environments in which pest and disease are likely to become established.
  4. As a minimum this will require the AA site to implement and keep associated records of a periodic inspection regime and ensure knock-down spray (such as standard household aerosol insecticide spray) is kept on-site at all times. A document outlining all pest control measures must be available to the department for audit purposes (example attached).
  5. In addition to details of the inspection regime and the on-site location of the knock-down spray, this document may include:
* the use of insecticides, fumigation, rodenticides, periodic inspection, baits and/or traps
* site plan with numbered bait stations
* contract details if applicable.

Note: The operations of adjacent facilities must be considered when determining any additional pest control measures to be implemented.

### Biosecurity area

* 1. The biosecurity area must be of a size commensurate with the proposed quantity of goods being handled.
  2. Biosecurity areas must be managed to allow biosecurity officers to conduct adequate inspections of goods in a timely and effective manner. This can occur by having:
* illumination to a sufficient level (within a building this will require a minimum 400 lux in storage areas and 600 lux in biosecurity inspection areas)
* goods accessible for inspection (this will require that goods be stored no more than 2.5 m high unless racks are used).

Note: ‘Accessible’ means goods must be able to be inspected as directed by a biosecurity officer. Generally, block stacking will not be regarded as being accessible.

### Security

* 1. Biosecurity areas where goods subject to biosecurity control are stored or handled must display a biosecurity sign to assist in effectively managing the security of these goods. These signs are to be:
* secured on a building/s, racks, fences, gates and/or doors and be visible at all times
* permanently affixed
* of a professional standard
* made to state ‘Biosecurity Area – Authorised Persons Only, No Entry or Removal of Goods, Penalties Apply, Biosecurity Act 2015’ or ‘Quarantine area – Authorised Persons Only, No Entry or Removal of Goods, Penalties Apply, ‘Quarantine Act 1908’or as directed for specific biosecurity operations
* on a yellow background with black lettering.

Note 1*:* Cardboard and paper signs are not acceptable.

Note 2: Where new signs are being produced, they should use ‘biosecurity’ not ‘quarantine’.

* 1. Signs on external structures must be:
* a minimum 600 mm x 400 mm with lettering a minimum 25 mm height
* weatherproof and resistant to the elements

Signs within structures must be a minimum 295 mm x 210 mm with lettering a minimum 8 mm height (example attached).

* 1. The following procedures must be applied to manage the AA site in a way that effectively secures goods subject to biosecurity control from movement or interference by unauthorised persons:
* the department must be immediately informed of any incidents which could significantly compromise the biosecurity integrity of the facility. This may include structural damage, electrical breakdowns, escapes or unauthorised entry and the removal of material subject to biosecurity control
* goods subject to biosecurity control must be stored in an area that is securely locked when unattended.

Note: Video surveillance, alarms or other security monitoring methods may also be used.

### Operating procedures

* 1. A document detailing procedures for the clean-up of biosecurity related spills must be available to the department for audit purposes. This document must include:
* the equipment used
* where applicable, the cleaning of this equipment (via disinfectant, sterilisation, or other approved method) and the spillage area with an approved broad-spectrum disinfectant.
  1. Biosecurity related spills include any spillage of goods subject to biosecurity control, waste or waste water. These spills must be disposed of in a manner as per the section on biosecurity waste. Equipment used for the clean-up of biosecurity related spills must be provided. Broad-spectrum disinfectants can be found on the department’s website.
  2. Any major spillage or loss of material subject to biosecurity control must be immediately reported to the department.

Note: A major spillage is classified as a loss of material subject to biosecurity control outside the confines of the AA site, which cannot be readily cleaned up within 15 minutes, or which may be accessed by the general public.

* 1. The BIP must provide a document detailing the entire imported goods pathway. This document should include all biosecurity operations.
  2. A procedure must be in place that ensures that the department is notified of any pest or disease infestation.

### Administration and management

* 1. Recordkeeping procedures must provide the department with the confidence that the system has adequate controls and the necessary evidence to verify identifiable links to goods. This can be achieved by:
* electronic or manual records of goods subject to biosecurity control imported through the AA site. This includes retaining originals or copies of Import Permits, biosecurity entries/directions or transfer approvals
* retaining records for a minimum period of 18 months after release from biosecurity control or disposal of the goods
* ensuring that records are available within 48 hours for inspection by the department.

Note: The department will continue to assess whether activities and arrangements have been implemented effectively and are meeting AA requirements. If records are unavailable during an inspection/audit, the department will return to the AA site within 48 hours to continue the assessment of documentation. This inspection will be charged at fee-for-service rates.

* 1. Office and general AA site requirements must provide the department with the confidence that applicable work health and safety standards have been met, this is achieved by:
* providing a first aid cabinet/kit which is fully stocked and meets the minimum commercial Australian Standard (AS2675-1983: Portable first aid kits for use by consumers)
* providing vehicle parking for visiting biosecurity officers. Note: this may require Department of Agriculture and Water Resources identified parking or providing a parking permit
* ensuring adequate security for any Department of Agriculture and Water Resources technical equipment left on the AA site
* providing access to and the availability of:
* a desk, chair and a telephone with direct outside call access
* toilet facilities
* handwashing facilities and a hygienic means of drying hands
* suitable arrangements to ensure amenities are clean.

Note: Additional mandatory requirements apply to AA sites with permanent biosecurity officers.

* 1. The AA site must comply with all relevant safety codes and work health and safety legislation.
  2. Administration and documentation requirements must provide the department with assurance that there are adequate controls. This must include:
* applications being accompanied by scale drawings (with dimensions and locations of biosecurity areas), identifying facilities for treatments, nearest main road and parking for biosecurity officers
* obtaining a departmental direction or prior written approval to move, accept, transfer or release any goods subject to biosecurity control from the approved facility to another approved facility that is not co-located
* where applicable, developing a transfer procedure for the safe movement of goods subject to biosecurity control between co-located facilities. This procedure must be provided at application and at the request of a biosecurity officer.
  1. BIP will need to apply in writing requesting authority to transfer goods subject to biosecurity control to an AA site not co-located when a direction, written approval or an applicable Import Permit has not been issued. This will require details of proposed suitable transport containers if applicable, the intended transport route and any other relevant information to support the case. The department may seek further information before making a decision.
  2. Control and security of the biosecurity area is the responsibility of the BIP. The BIP must ensure that there is an agreement with the department for training and/or electronic initiatives needed. Failure to comply with the approval requirements or any violation of the Act may result in the approval of the AA site being withdrawn or suspended and legal action prompted.
  3. The BIP must ensure that persons having physical access to goods subject to biosecurity control are aware that such items must only be handled by an Accredited Person or under the direct supervision of an Accredited Person.
  4. Persons performing the function of an Accredited Person must have successfully completed the department's approved training to obtain and maintain Accredited Person status.
  5. Records must be maintained of Accredited Persons.
  6. The BIP must notify the department in writing as soon as practicable within 15 working days of becoming aware of any change of status, not previously been notified to the department, of the BIP or their associates relevant to the operation of the AA in relation to any of the following matters:
* conviction of an offence or order to pay a pecuniary penalty under the Biosecurity Act 2015, Quarantine Act 1908, Customs Act 1901, the Criminal Code or the Crimes Act 1914
* debt to the to the Commonwealth that is more than 28 days overdue under the Biosecurity Act 2015, Quarantine Act 1908, Customs Act 1901, the Criminal Code or the Crimes Act 1914
* refusal, involuntary suspension, involuntary revocation/cancelation or involuntary variation of an Import Permit, quarantine approved premises, compliance agreement or AA under the Quarantine Act 1908 or the Biosecurity Act 2015.
  1. Biosecurity officers, biosecurity enforcement officers and department-approved auditors, must be provided access to the AA site to perform the functions and exercise the powers conferred on them by the Biosecurity Act or another law of the Commonwealth.
  2. Departmental auditors or department-approved auditors must be provided with facilities and assistance as requested, and any required documents, records or things relevant to the audit.
  3. The department must be notified of any Reportable Biosecurity Incident as soon as practicable, in accordance with the determination made by the Director of Biosecurity.
  4. Department-approved auditors must be permitted to collect evidence of compliance and noncompliance with AA requirements through actions including the copying of documents and taking of photographs.

## Specific requirements

This part outlines the specific requirements that must be complied with for facilities to be approved to a specific facility type.

Note: These requirements are additional to 3.1: About BC2 requirements and the requirements for approval and Part 3: General requirements.

### Microbiological containment—level 3 (BC3) facilities

### Hygiene and isolation

* 1. Write-up areas may be approved as part of a BC3 facility. To be eligible for approval, these areas must comply with BC3 requirements, be constructed such that horizontal surfaces are minimised (for example, minimal shelving), must not be used for generic office functions and should hold only essential reference materials (such as technical equipment manuals).

### Waste disposal

* 1. Where applicable any biosecurity waste must be effectively contained and disposed of in a manner approved by the department and be detailed in a document outlining specific procedures for the disposal of any accumulated waste which may include a section on the:
* disposal of waste that is not subject to Import Permit conditions
* movement of waste within the AA site where an approved method is not available within the biosecurity area/facility.
  1. Solid biosecurity waste must be bagged and placed in an unbreakable container with a secured lid for movement within the building to the approved disposal place.
  2. Where waste cannot be disposed of immediately, there must be, at minimum, the provision for:
* a separate storage device/area for the temporary holding of goods
* storage in lidded bins/containers of an appropriate size which are leak and pest-proof
* bins to be labelled ‘Biosecurity Waste’
* double bagging of waste.
  1. Separate storage device/area must be approved by the department and be within the AA site to prevent loss, spillage or unauthorised access.
  2. Approved methods of solid biosecurity waste disposal include incineration at a high temperature in a high efficiency EPA approved incineration facility, deep burial, or sterilisation by autoclaving.
  3. Minimum autoclaving times after attainment of temperature for goods, residues or biosecurity waste shall be either 121°C:
* (core temperature) for 15 minutes
* for 30 minutes where core temperature is not measured.
  1. Where the 15 minute autoclaving time is used, the BIP must specify how the core temperature has been reached and detail how this temperature was recorded.
  2. Biosecurity waste water must be disposed of by an approved method.

Note: This will require a system to be in place which decontaminates the effluent from the biosecurity/work area, inner change room and, where applicable, shower before being discharged. An alarm must be provided to alert persons of any decontamination system malfunction.

* 1. The use of a waste water disposal method must be approved in writing by the department and may require demonstration of their efficacy to the satisfaction of the department. Such methods may require detailed scientific research at the BIP’s expense.

### Security

* 1. To assist in effectively managing the security of the facility the following must be applied:
* a logbook kept, recording visitor names, their company and the time and date of visits
* the name and telephone number of the facility manager or other responsible person must be displayed near the access doors
* a biosecurity sign be displayed on the entry door to the facility. Such signs are to include the requirements as stated in 6.1, and in addition, state ‘Microbiological Containment – BC3 (or QC3) Facility’.

Note: Where new signs are being produced, they should use ‘BC’ not ‘QC’.

* 1. Access to the facility for cleaning, servicing of equipment and repairs must not occur until potentially contaminated surfaces have been disinfected with a department-approved broad-spectrum disinfectant.

### Operational procedures

* 1. Containers holding goods subject to biosecurity control must be clearly labelled using standard scientific nomenclature. The labelling must enable clear reconciling of goods subject to biosecurity control with the following information:
* biosecurity entry number (where relevant)
* Import Permit number or departmental in vivo approval number and expiry dates
* importation date.

If the containers cannot be labelled with this information due to constraints, such as size, then a suitable identification system may be used, such as referring to a logbook that contains the required information.

* 1. Equipment used or that has come into contact with goods subject to biosecurity control must be cleaned or rendered safe by an approved method. Department-approved methods include, but are not limited to:
* sterilisation
* incineration, as prescribed in part 10: Waste disposal
* disinfection using a department-approved broad-spectrum disinfectant.
  1. Equipment must be disinfected before being sent for repair or disposal and, where appropriate, equipment may need to be disinfected at regular intervals.
  2. Gloves must be removed and hands thoroughly washed with soap and warm water after handling goods subject to biosecurity control, working in a biological safety cabinet and before leaving the facility. Used gloves shall be discarded with the biosecurity waste.
  3. To prevent cross-contamination while work is being undertaken there must be separation of biosecurity work from other work.
  4. Prior to entering the facility and in the anteroom, personnel must put on covering clothes. Additionally, these garments must be:
* removed on leaving the facility and kept in the anteroom (or laboratory) between uses
* laundered at appropriate intervals
* where disposable protective clothing is used it must be disposed of in the manner described in the waste disposal section of this document.

Note: The department must be provided with a written procedure of how protective clothing will be laundered.

* 1. If a separate culture room within the facility exists and is dedicated to this work, then complete protective clothing including overshoes and hats is required within this room. Dedicated apparel should be used in this culture room and not removed except for laundering or disposal.
  2. When not in use, containers of regulated articles must be stored securely in the biosecurity area (for example, in cool rooms, incubators, refrigerators, cupboards or similar structures).
  3. A document outlining an inspection regime for all goods must be provided to the department for audit purposes. The minimum requirements for this document include the interval and personnel who conducted the inspection.
  4. To ensure the containment features of the facility are intact a document detailing annual inspection must be available to the department for audit purposes. The minimum requirements for this document include:
* an annual inspection report detailing findings, including HEPA filter integrity test reports and room pressure readings
* the personnel who conducted the inspection.

Note: Personnel include the BIP, employees or the Biosafety Committee.

* 1. A negative pressure gradient shall be maintained that exists of 25 Pa in the anteroom and 50 Pa in the AA site relative to the pressure outside the facility. Additionally, the pressure inside the facility shall never fall below 25 Pa, relative to the pressure outside of the facility, when the door between the anteroom and the AA site is open.
  2. Annual testing and certification by a qualified technician must include:

1. testing of the pressure differentials in accordance with AS 1807.10
2. integrity testing of all installed HEPA filters in accordance with AS 1807.6 or AS 1807.7
3. checking that the control system is operating correctly and verifying alarms are set to operate when room differential air pressures depart from set points by more than 15 Pa for a period of greater than two minutes
4. the effectiveness of the effluent treatment and decontamination system. Note: any failures of the effluent treatment system must be rectified and the system re-tested. Records must be retained
5. a report of the testing in items 1) to 4) and of any maintenance conducted must be provided at the request of a biosecurity officer.
   1. A manual listing procedures and documents applicable to biosecurity, including emergency and maintenance procedures, must be available within the facility in a prominent position.
   2. The BIP must provide the department with information concerning the calibration and certification of pressure steam sterilisers and the efficacy of treatment. The minimum requirements for sterilisers are that:

* relevant local regulations for pressure vessels are applied, including the timeframes for the regular certification of the steriliser.
* steriliser cycles are calibrated. This can be achieved by the use of either:
* thermocouples or resistance thermometers to ensure that the sterilisation temperature indicated has been achieved
* chemical indicators which progressively change colour with the time exposed at the specified temperature
* biological indicators such as spore strips
* enzyme indicators be used at regular intervals (for example, monthly)
* other approved methods.

Note:

* the timing of the sterilisation stage of the cycle commences when the set temperature is recorded by the thermometer in the drain line
* where indicators are used they must be placed in several positions in a load, including those least likely to attain sterilization conditions
* sterilisers must be located in the facility or within the building that houses the facility.
  1. The annual checking and certification of sterilisers or heat ovens must be carried out by a qualified technician.
  2. The BIP must provide the department with information concerning the efficiency and safety of biological safety cabinets. The minimum requirements for biological safety cabinets are that:
* cabinets are checked for containment efficiency and safety before initial use, after any modification including change of HEPA filters, after relocation and on an annual basis
* used filters are disposed of with biosecurity waste
* where class II cabinets are used they will need to pass the air barrier containment test in AS 1807.22.

Note: The annual checking and certification of cabinets must be carried out by a qualified technician.

* 1. Flexible film isolators must operate at a pressure below that of the room in which they are located and inlet and exhaust systems must be HEPA filtered. The annual checking and certification of isolators must be carried out by a qualified technician.

### Administration and management

* 1. Records for each consignment of goods subject to biosecurity control must include:
* biosecurity entry number (where relevant)
* Import Permit number or departmental in vivo approval number for the regulated articles
* description of the regulated goods (using accurate scientific terminology)
* date of receipt of goods and country of origin
* location or part of facility where each item subject to biosecurity control is held, and the respective BC status
* records of any derivatives and additional cultures/material or substance grown from the original material subject to biosecurity control
* where applicable quantities (for example, kg or litres) of goods received, destroyed and in storage
* date of completion of research
* details of any treatments
* method and date of disposal/destruction of goods subject to biosecurity control and any direct or indirect derivatives
* method, and date of waste disposal/destruction
* the date of movement and the department’s permission for any movement (including transfer certificates) of goods from the facility
* comprehensive details of any breaches of goods subject to biosecurity control from the facility.
  1. A bi-annual summary of records, that includes the information in 13.1, must be provided at audit or at the request of a biosecurity officer.
  2. A logbook (electronic or manual) recording steriliser load, temperature and duration of sterilisation cycle must be maintained.
  3. Calibration specifications for equipment that has a bearing on the biosecurity status of the material (such as autoclave), along with calibration records must be provided at audit and at the request of a biosecurity officer.
  4. Drinking water, food and toilets shall only be located within designated areas of the biosecurity facility where goods are not handled, stored or treated and there is also no potential for harm to animals, plants, humans or the environment. A procedure must be in place which ensures gloves and garments are removed and hands are washed prior to drinking, eating or use of the toilet facilities.

### Indoor animal containment—level 3 (BC3) facilities

### Isolation and hygiene

* 1. Where post mortem examinations are undertaken the following conditions apply:
* a separate area from other activities such as animal production must be provided
* adequate precautions taken to prevent cross-contamination.
  1. Secure housing/caging must be provided.
  2. Where the facility acts as the primary containment, a body shower and inner and outer change rooms must be incorporated.
  3. Write-up areas may be approved as part of a BC3 facility. To be eligible for approval, these areas must comply with BC3 requirements, be constructed such that horizontal surfaces are minimised (for example, minimal shelving), must not be used for generic office functions and should hold only essential reference materials (such as technical equipment manuals).

### Waste disposal

* 1. A document must be provided to the department outlining how carcasses from animals subject to biosecurity control will be effectively contained and rendered safe prior to disposal in a manner approved by the department.
  2. This document should cover specific procedures for the disposal of any carcasses. This may include transportation (where the carcass has not been rendered safe at the AA site) by an approved transporter, or under the department’s supervision.
  3. Procedures where carcasses cannot be disposed of immediately should also be covered. This may include the provision for a separate storage device/area. Such areas and/or devices must be insect, rodent and bird proof.
  4. The separate storage device/area must be the department approved and be within the AA site to prevent loss or unauthorised access.
  5. Approved methods of disposal of carcass subject to biosecurity control include incineration at a high temperature in a high efficiency EPA approved incineration facility, deep burial, or sterilisation by autoclaving.
  6. Minimum autoclaving times after attainment of temperature for the goods, residues or biosecurity waste shall be either 121°C:
* (core temperature) for 15 minutes
* for 30 minutes where core temperature is not measured.
  1. Where the 15 minute autoclaving time is used, the BIP must specify how the core temperature has been reached and detail how this temperature was recorded.
  2. Animal bedding must be disposed of by an approved method. Approved methods include, but are not limited to, incineration at a high temperature in a high efficiency EPA-approved incineration facility, sterilisation or deep burial.
  3. Provision must be made for the decontamination of pens and cages. Decontamination can be achieved by:
* using an approved broad-spectrum disinfectant
* an approved method.
  1. Biosecurity waste water must be disposed of by an approved method. This will require a system to be in place which decontaminates the effluent from the biosecurity/work area, inner change room and, where applicable, shower before being discharged.
  2. An alarm must be provided to alert persons of any decontamination system malfunction. The use of a waste water disposal method must be approved in writing by the department and may require demonstration of their efficacy to the satisfaction of the department. Such methods may require detailed scientific research at the BIP’s expense.
  3. Where a facility performs a primary containment function and animal holding areas/pens/cages are plumbed to floor drains, these drains must be fitted with traps to ensure that solids (such as bedding, faecal matter) are collected during research and at times of pen/cage washing and disinfection. Waste solids collected from drains must be treated by an approved method.

Approved methods of solid biosecurity waste disposal include incineration at a high temperature in a high efficiency EPA approved incineration facility, deep burial, or sterilisation by autoclaving.

* 1. Where the facility has floor drains, the drain traps should always be filled with water and a suitable approved broad spectrum disinfectant, and be secure against entry by pests.

### Security

* 1. A nominated staff member employed by the AA site is responsible for gate, door or animal enclosure keys. The name, designation/position title and contact details of this person must be supplied with the application form and when a change in the nominated staff member occurs.
  2. To assist in effectively managing the security of the facility the following must be applied:
* a logbook kept, recording visitor names, their company and the time and date of visits
* the name and telephone number of the facility manager or other responsible person must be displayed near all access doors
* a biosecurity sign is displayed on the entry door to the facility. Signs on the entry door to the facility must include the requirements as stated in 6.1, and in addition, state ‘Animal containment – BC3 (or QC3) facility’.

Note: Where new signs are being produced, they should use ‘BC’ not ‘QC’.

* 1. Access to the facility for cleaning, servicing of equipment and repairs must not occur until potentially contaminated surfaces have been disinfected with an approved broad-spectrum disinfectant.

### Operational procedures

* 1. Arrangements must be in place for animals undergoing in vivo trials involving imported biologicals that ensures daily checking. A written record must be kept of daily checks.
  2. Identification must be possible for animals subject to biosecurity control (for example, by tattooing, microchip, permanent branding or through a cage labelling system).
  3. Where applicable, cages and racks must be labelled to indicate the identity and date of any inocula given.
  4. Where it is necessary to transport animals (alive or dead) from the containment facility, in addition to the requirements in 8, Administration, procedures must include details on pens/cages used for transport and the decontamination of these with the department’s approved broad-spectrum disinfectant.
  5. Unexpected animal mortalities or incidence of disease must be reported to the department immediately and investigated. This may require instructions regarding:
* the animals being labelled with day/date
* where possible preserved (in a refrigerator, coolroom or freezer) for appropriate post mortem and examination by a biosecurity officer or a suitably qualified veterinarian employed by the BIP.
  1. In the case where the investigation is conducted by the BIP, the department must be kept informed on the progress of the investigation, and must be provided a report at the conclusion of the investigation.
  2. Gloves shall be removed and hands thoroughly washed with soap and warm water after handling goods subject to biosecurity control, working in a biological safety cabinet and before leaving the facility. Used gloves shall be discarded with the biosecurity waste.
  3. Where the facility acts as the primary containment barrier, personnel must leave the facility through the clothing, change and shower rooms, except in cases of emergency where alternative exits may be used.
  4. Street clothing, including underwear, shall be removed and retained in the outer clothing change room. Complete protective clothing, including shoes must be used by the personnel entering the facility.
  5. When leaving the facility, personnel must remove their laboratory clothing and store or discard it in the inner change room before showering.
  6. Where disposable protective clothing is used it must be disposed of in the manner described in the waste disposal section of this document. Other clothing must be laundered at appropriate intervals.
  7. Prior to entering the facility and in the anteroom, or where applicable, inner change room, the following minimum requirements apply to clothing and other apparel personnel must:
  8. Additionally, these garments must be:
* removed on leaving the facility and kept in the anteroom between uses
* laundered at appropriate intervals
* where disposable protective clothing is used it must be disposed of in the manner described in waste disposal, of this document.

Note: The department must be provided with a written procedure of how covering clothes will be laundered

* 1. Where the facility has floor drains, the drain traps should always be filled with water and a suitable approved broad spectrum disinfectant, and be secure against entry by pests.
  2. A negative pressure gradient shall be maintained that exists of 25 Pa in the anteroom and 50 Pa in the AA site relative to the pressure outside the facility. Additionally, the pressure inside the facility shall never fall below 25 Pa, relative to the pressure outside of the facility, when the door between the anteroom and the AA site is open.
  3. Annual testing and certification by a qualified technician must include:

1. Testing of the pressure differentials in accordance with AS 1807.10.
2. Integrity testing of all installed HEPA filters in accordance with AS 1807.6 or AS 1807.7.
3. Checking that the control system is operating correctly and verifying alarms are set to operate when room differential air pressures depart from set points by more than 15 Pa for a period of greater than 2 minutes.
4. The effectiveness of the effluent treatment and decontamination system

Note: Any failures of the effluent treatment system must be rectified and the system re-tested. Records must be retained.

1. A report of the testing in items 1) to 4) and of any maintenance conducted must be provided at the request of a biosecurity officer.
   1. A manual listing procedures and documents applicable to biosecurity, including emergency and maintenance procedures must be available within the facility in a prominent position.
   2. Pressure steam sterilisers must be located in the facility or within the building that houses the facility.
   3. The BIP must provide the department with information concerning the calibration and certification of sterilisers, and the efficacy of treatment. The minimum requirements for sterilisers are that:

* relevant local regulations for pressure vessels are applied, including the timeframes for the regular certification of the steriliser.
* steriliser cycles are calibrated. This can be achieved by the use of either:
* thermocouples or resistance thermometers to ensure that the sterilisation temperature indicated has been achieved
* chemical indicators which progressively change colour with the time exposed at the specified temperature
* biological indicators such as spore strips
* bacterial enzyme indicators are used at regular intervals (e.g. monthly)
* other approved method.

Note: The timing of the sterilisation stage of the cycle commences when the set temperature is recorded by the thermometer in the drain line.

Where indicators are used they must be placed in several positions in a load, including those least likely to attain sterilisation conditions.

* 1. The annual checking and certification of sterilisers or heat ovens must be carried out by a qualified technician.

### Administration and management

* 1. Records for each consignment of goods subject to biosecurity control must include:
* date of receipt of goods and country of origin
* Import Permit number or in vivo approval
* location or part of facility where each item subject to biosecurity control is held
* date of completion of research
* details of any treatments
* method and date of goods disposal/destruction (if applicable), and any direct or indirect derivatives
* the date and the department’s permission for any movement (including transfer certificates) of goods from the facility
* comprehensive details of any breaches of goods subject to biosecurity control from the facility.
  1. A record must be maintained of an up to date inventory of the animals present and a chronological record of procedures performed.
  2. Records should be kept of births, (if applicable), mortalities, post mortem findings and test results.
  3. Details of post mortem results must be made available at the request of a biosecurity officer.
  4. Drinking water, food and toilets shall only be located within designated areas of the biosecurity facility where goods are not handled, stored or treated and there is also no potential for harm to animals, plants, humans or the environment. A procedure must be in place which ensures gloves and garments are removed and hands are washed prior to drinking, eating or use of the toilet facilities.

### Plant laboratory containment—level 3 (BC3) facilities

### Hygiene and isolation

* 1. Write-up areas may be approved as part of a BC3 facility. To be eligible for approval, these areas must comply with BC3 requirements, be constructed such that horizontal surfaces are minimised (for example, minimal shelving), must not be used for generic office functions and should hold only essential reference materials (such as technical equipment manuals).

### Waste disposal

* 1. A document must be provided to the department describing how biosecurity waste should be effectively contained and rendered safe prior to disposal in a department-approved manner.
  2. This document should cover specific procedures for the removal of any accumulated waste. This may include:
* that which is not subject to Import Permit conditions
* movement within the AA site where an approved method is not available within the biosecurity area/facility.
  1. Solid biosecurity wastes must be bagged and placed in an unbreakable container with a secured lid for movement within the building to the approved disposal place.
  2. Procedures where waste cannot be disposed of immediately should also be covered. This must as a minimum include the provision for:
* a separate storage device/area for the temporary holding of goods
* storage in lidded bins/containers of an appropriate size which are leak and pest proof
* bins to be labelled ‘Biosecurity Waste’
* the waste must be double bagged.
  1. The separate storage device/area must be approved by the department and be within the facility to prevent loss, spillage or unauthorised access.
  2. Approved methods of solid biosecurity waste disposal include incineration at a high temperature, in a high efficiency EPA approved incineration facility, deep burial or sterilisation by autoclaving.
  3. Minimum autoclaving times after attainment of temperature for the goods, residues or biosecurity waste shall be either 121°C:
* (core temperature) for 15 minutes
* for 30 minutes where core temperature is not measured.
  1. Where the 15 minute autoclaving time is used the BIP must specify how the core temperature has been reached and detail how this temperature was recorded.
  2. Biosecurity waste water must be disposed of by a department-approved method. This will require a system to be in place which decontaminates the effluent from the biosecurity/work area, inner change room and, where applicable, shower before being discharged.
  3. An alarm must be provided to alert persons of any decontamination system malfunction.
  4. The use of a waste water disposal method must be approved in writing by the department and may require demonstration of their efficacy to the satisfaction of the department. Such methods may require detailed scientific research at the BIP’s expense.

### Security

* 1. To assist in effectively managing the security of the facility the following must be applied:
* the name and telephone number of the facility manager or other responsible person must be displayed near the access doors
* a biosecurity sign be displayed on the entry door to the facility. Signs on the entry door to the facility must include the requirements as stated in 6.1, and in addition, state ‘Plant containment – BC3 (of QC3) facility’.

Note: Where new signs are being produced, they should use ‘BC’ not ‘QC’.

### Operational procedures

* 1. To ensure the containment features of the facility are intact a document detailing annual inspection must be available to the department for audit purposes. The minimum requirements for this document include:
* an annual inspection report detailing findings
* the personnel who conducted the inspection.

Note: Personnel include the BIP, employees or the Biosafety Committee.

* 1. The BIP must provide documentary evidence that screens, filters and similar equipment have been cleaned in accordance with the manufacturer’s specified frequency and procedures. This can be achieved by:
* supplying the frequency plan and procedures provided by the manufacturer
* recording the date that the cleaning occurred.
  1. Unexpected incidences of pest or disease must be reported to the department immediately.
  2. Gloves shall be removed and hands thoroughly washed with soap and warm water after handling goods subject to biosecurity control, working in a biological safety cabinet and before leaving the facility. Used gloves shall be discarded with the biosecurity waste.
  3. Prior to entering the facility and in the anteroom, the following minimum requirements apply to clothing personnel must put on covering clothes.
  4. Additionally, these garments must be:
* removed on leaving the facility and kept in the anteroom (or laboratory) between uses
* laundered at appropriate intervals
* where disposable protective clothing is used it must be disposed of in the manner described in waste disposal, of this document.

Note: The department must be provided with a written procedure of how protective clothing will be laundered.

* 1. If a separate culture room within the facility exists and is dedicated to this work, then complete protective clothing including overshoes and hats is required within this room. Dedicated apparel should be used in this culture room and not removed except for laundering or disposal.
  2. A negative pressure gradient shall be maintained that exists of 25 Pa in the anteroom and 50 Pa in the AA site relative to the pressure outside the facility. Additionally, the pressure inside the facility shall never fall below 25 Pa, relative to the pressure outside of the facility, when the door between the anteroom and the AA site is open.
  3. Annual testing and certification by a qualified technician must include:

1. Testing of the pressure differentials in accordance with AS 1807.10.
2. Integrity testing of all installed HEPA filters in accordance with AS 1807.6 or AS 1807.7.
3. Checking that the control system is operating correctly and verifying alarms are set to operate when room differential air pressures depart from set points by more than 15 Pa for a period of greater than 2 minutes.
4. The effectiveness of the effluent treatment and decontamination system.

Note: Any failures of the effluent treatment system must be rectified and the system re-tested. Records must be retained.

1. A report of the testing in items 1) to 4) and of any maintenance conducted must be provided at the request of a biosecurity officer.
   1. A manual listing procedures and documents applicable to biosecurity including emergency and maintenance procedures must be available within the facility in a prominent position.
   2. Pressure steam sterilisers must be located in the facility or within the building that houses the facility.
   3. The BIP must provide the department with information concerning the calibration and certification of sterilisers, and the efficacy of treatment. The minimum requirements for sterilisers are that:

* relevant local regulations for pressure vessels are applied, including the timeframes for the regular certification of the steriliser
* steriliser cycles are calibrated. This can be achieved by the use of either:
* thermocouples or resistance thermometers to ensure that the sterilisation temperature indicated has been achieved
* chemical indicators which progressively change colour with the time exposed at the specified temperature
* biological indicators such as spore strips
* bacterial enzyme indicators be used at regular intervals (e.g. monthly)
* other departmental approved methods.

Note: The timing of the sterilisation stage of the cycle commences when the set temperature is recorded by the thermometer in the drain line. Where indicators are used they must be placed in several positions in a load, including those least likely to attain sterilisation conditions.

* 1. The annual checking and certification of sterilisers or heat ovens must be carried out by a qualified technician.
  2. The BIP must provide the department with information concerning the efficiency and safety of biological safety cabinets. The minimum requirements for biological safety cabinets are that:
* cabinets be checked for containment efficiency and safety before initial use, after any modification including change of HEPA filters, after relocation and on an annual basis
* used filters be disposed of with biosecurity waste
* where class II cabinets are used they will need to pass the air barrier containment test in AS 1807.22.

Note: The annual checking and certification of cabinets must be carried out by a qualified technician.

* 1. Flexible film isolators must operate at a pressure below that of the room in which they are located and inlet and exhaust systems are HEPA filtered. Annual checking and certification of isolators must be carried out by a qualified technician.

### Administration and management

* 1. Records for each consignment of goods subject to biosecurity control must include:
* date of receipt of goods and country of origin
* Import Permit number, in vivo approval number or transfer approval
* plant material type (where applicable include scientific name)
* location or part of facility where each item subject to biosecurity control is held
* date of completion of research
* details of any treatments
* method and date of disposal/destruction of goods subject to biosecurity control (if applicable) and any direct or indirect derivatives
* the date of movement and the department’s permission for any movement (including transfer certificates) of goods from the facility
* comprehensive details of any breaches of goods subject to biosecurity control from the facility.
  1. A record must be maintained of an up to date inventory of the plant material present and a chronological record of procedures performed.
  2. A logbook (electronic or manual) recording steriliser load, temperature and duration of sterilisation cycle must be maintained.
  3. Drinking water, food and toilets shall only be located within designated areas of the biosecurity facility where goods are not handled, stored or treated and there is also no potential for harm to animals, plants, humans or the environment. A procedure must be in place which ensures gloves and garments are removed and hands are washed prior to drinking, eating or use of the toilet facilities.

## Applicable Australian/New Zealand Standards

This part outlines the specific standards that a department-approved ‘third party’ assessor will certify. The applicable parts of the standards must be complied with for facilities to be approved to a specific facility type.

Where reference is made to an Australian/New Zealand Standard (or clause in an Australian/New Zealand Standard) in the requirements against which a facility is to be certified, that referenced standard (or clause) must also be met.

Note: These requirements are additional to Part 3.1: About BC3 requirements and the requirements for approval, Part 3: General requirements, and Part 4: Specific Department of Agriculture and Water Resources requirements for BC3 approval of a particular facility type.

Australian/New Zealand Standard – Laboratory Design and Construction Part 1: General Requirements (AS/NZS 2982.1:1997)

The following structural parts of this standard (AS/NZS 2982.1:1997) are applicable to BC Level 3.

### Specific standards

### Microbiological containment—level 3 (BC3) facilities

* 1. The BIP of microbiological containment approval must also meet Part 3.1 about the BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for Microbiological containment.
  2. The following standards from AS/NZS 2982.1:1997 are the minimum for work with microbiological goods at the BC3 level.

Section 2: General laboratory design and construction requirements (excluding 2.2, 2.6, 2.7 (a) (vi) & (vii) and 2.7 (c), 2.9, 2.11, 2.12 and 2.13).

Section 3: Reticulated services (excluding 3.7.3).

Section 4: Electrical services (excluding 4.2 and 4.3).

Section 5: Ventilation and air quality (excluding 5.1, 5.2, 5.3, 5.5.1, 5.5.3, 5.6 (b) last sentence only and 5.7).

Section 6: Health and safety requirements (excluding 6.1, 6.4, 6.5 and 6.6).

* 1. In substitution for 6.2 (safety showers), the following clause will be applied:

Clean up provisions are required which may be either:

* fixed appliances (for example, showers and eyewash stations)
* single use apparatus (for example, disinfectant swabs, squeeze bottles).

Note: Where safety showers and fixed eyewash facilities are in use they must comply with the requirements of 6.2. The approach to clean up provisions must be unobstructed.

* 1. In substitution for 6.3 (hand washing facilities) the following clause will be applied:

Work areas where goods subject to biosecurity control are handled must contain either a handwash basin fitted with hands-free taps, or some other means of decontaminating hands.

Note: Handwash basins must be located inside the laboratory, near to the exit serviced with hot and cold potable water. Potable water requirements must be met in accordance with AS 3500.

* 1. Alternatives to wash basins include:
* dispensers fitted with approved antiseptic solutions, are considered suitable, provided the dispensers can be operated without using the hands
* a sink of hands-free operation.

Section 8: Biological laboratories (only 8.3).

Appendix B: Additional requirements for Microbiological facilities (only B4 (f) and B5 (i) iv).

* 1. In addition to the above standards where a basin is provided for washing hands; an antiseptic handwash dispenser must be supplied.

### Indoor animal containment—level 3 (BC3) facilities

* 1. The BIP of indoor animal containment approval must also meet Part 3.1 about BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for Indoor animal containment.
  2. The following standards from AS/NZS 2982.1:1997 are the minimum for work with indoor animal goods at the BC3 level.

Section 2: General laboratory construction requirements (excluding 2.2, 2.6, 2.7 (a) (vi) & (vii) & 2.7 (c); 2.9, 2.10, 2.11, 2.12 and 2.13).

Section 3: Reticulated services (excluding 3.7.3).

Section 4: Electrical services (excluding 4.2 and 4.3).

Section 5: Ventilation and air quality (excluding 5.1, 5.2, 5.3, 5.5.1, 5.5.3, 5.6 (b) last sentence only and 5.7).

Section 6: Health and safety requirements (excluding 6.1, 6.4, 6.5 and 6.6).

* 1. In substitution for 6.2 (safety showers), the following clause will be applied. Clean up provisions are required which may be either:
* fixed appliances (e.g. showers and eyewash stations)
* single use apparatus (for example, disinfectant swabs, squeeze bottles).

Note: Where safety showers and fixed eyewash facilities are in use they must comply with the requirements of 6.2. The approach to clean up provisions must be unobstructed.

* 1. In substitution for 6.3 (hand washing facilities) the following clause will be applied:
  2. Alternatives to wash basins include:
* dispensers fitted with approved antiseptic solutions, are considered suitable, provided the dispensers can be operated without using the hands
* a sink of hands-free operation.

Section 8: Biological laboratories (only 8.6.8, 8.6.10 and 8.6.11).

Appendix C: Additional requirements for animal accommodation (only C3 (f)).

* 1. In addition to the above standards where a basin is provided for washing hands an antiseptic handwash dispenser be provided.

### Plant laboratory containment—level 3 (BC3) facilities

The BIP of plant laboratory containment approval must also meet Part 3.1 about BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for Plant Containment.

* 1. The following standards from AS/NZS 2982.1:1997 are the minimum for work with plant goods at the BC3 level.

Section 2: General laboratory construction requirements (excluding 2.2, 2.6, 2.7 (a) (vi) & (vii) & 2.7 (c); 2.9, 2.11, 2.12 and 2.13).

Section 3: Reticulated services (excluding 3.7.3).

Section 4: Electrical services (excluding 4.2 and 4.3).

Section 5: Ventilation and air quality (excluding 5.1, 5.2, 5.3, 5.5.1, 5.5.3, 5.6 (b) last sentence only and 5.7) .

Section 6: Health and safety requirements (excluding 6.1, 6.4, 6.5 and 6.6).

* 1. In substitution for 6.2 (safety showers), the following clause will be applied. Clean up provisions are required which may be:
* fixed appliances (such as showers and eyewash stations)
* single use apparatus (such as disinfectant swabs, squeeze bottles).

Note*:* Where safety showers and fixed eyewash facilities are in use they must comply with the requirements of 6.2. The approach to clean up provisions must be unobstructed.

* 1. In substitution for 6.3 (hand washing facilities) the following clause will be applied:

Work areas where goods subject to biosecurity control are handled must contain either a handwash basin fitted with hands-free taps, or some other means of decontaminating hands.

Note: Handwash basins must be located inside the laboratory, near to the exit serviced with hot and cold potable water. Potable water requirements must be met in accordance with AS 3500.

* 1. Alternatives to wash basins include:
* dispensers fitted with approved antiseptic solutions, are considered suitable, provided the dispensers can be operated without using the hands
* a sink of hands-free operation.

Section 8: Biological laboratories (only 8.3).

* 1. In addition to the above standards where a basin is provided for washing hands an antiseptic handwash dispenser be provided.

### Australian/New Zealand Standard—safety in laboratories: microbiological aspects and containment facilities (AS/NZS 2243.3:2002)

The following parts of this standard (AS/NZS 2243.3:2002) are applicable to BC3.

### Microbiological containment—level 3 (BC3) facilities

The BIP of microbiological containment approval must also meet Part 3.1 about the BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for Microbiological containment.

* 1. The following standards from AS/NZS 2243.3:2002 are the minimum for work with microbiological goods at the BC3 level.

Section 4.8: Physical containment level 2 (PC2) requirements

(Only 4.8.3 (b))

* 1. In addition to 4.8.3 (b) the following apply:

Large HVAC heat exchangers (e.g. chilled beams and some other natural or forced convectors) located within containment occupancies may be vulnerable to contamination, if they are not protected by efficient air filters. The likely contamination and implications for containment should be carefully assessed where such devices are proposed or are used.

Section 4.9: Physical containment level 3 (PC3) requirements (only 4.9.2 (a) – excluding the requirements for doors to open outwards, and be self-closing. The requirement for doors to open outwards will be assessed as flexible depending on the context of the facility and rooms within. 4.9.2 (b), (c) and 4.9.3 generic text and (a), (b), (c), (d), (e), (f) and (i)).

* 1. In substitution for 4.9.3 (g) the following clause will be applied. Each separate room within a facility must have:
* a room pressure gauge that can be viewed by personnel prior to entering the facility
* an audible alarm.
  1. Notes:
* the room pressure gauge must measure the differential pressure relative to the adjacent rooms external to the facility. The reference pressure should be taken from adjacent rooms, normally the access space not the ceiling space. The adjacent space should not be subject to significant variations in ambient static pressure due to external influences such as ambient winds, ventilation equipment, fans, or any other equipment that could give rise to pressure fluctuations
* the number of gauges required to safely monitor facility pressures will depend on the number of rooms, the number of independent supply and exhaust systems, and layout
* the audible alarm is to indicate loss of negative room pressure in excess of two minutes
* where practicable other air-conditioning control switches and exhaust fan speed set point control should also be located adjacent to the gauges
* the HEPA filter gauge can also be mounted with the room gauges
* if there is a closable door between two rooms they are separate.

### Indoor animal containment—level 3 (BC3) facilities

The BIP of indoor animal containment approval must also meet Part 3.1 about the BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for indoor animal containment.

* 1. The following standards from AS/NZS 2243.3:2002 are the minimum for work with indoor animal goods at the BC3 level

Section 4.8 PC2 requirements (only 4.8.3 (b)).

* 1. In addition to 4.8.3 (b) the following apply:

Large HVAC heat exchangers (e.g. chilled beams and some other natural or forced convectors) located within containment occupancies may be vulnerable to contamination, if they are not protected by efficient air filters. The likely contamination and implications for containment should be carefully assessed where such devices are proposed or are used.

Section 4.9: PC3 requirements (only 4.9.2 (a) – excluding the requirement for doors to open outwards, and be self-closing. The requirement for doors to open outwards will be assessed as flexible depending on the context of the facility and rooms within. 4.9.2 (b), (c) and 4.9.3 (a), (b), (c), (d), (e), (f) and (i))

* 1. In substitution for 4.9.3 (g) the following clause will be applied. Each separate room within a facility must have:
* a room pressure gauge that can be viewed by personnel prior to entering the facility
* an audible alarm.
  1. Notes:
* the room pressure gauge must measure the differential pressure relative to the adjacent rooms external to the facility. The reference pressure should be taken from adjacent rooms, normally the access space not the ceiling space. The adjacent space should not be subject to significant variations in ambient static pressure due to external influences such as ambient winds, ventilation equipment, fans, or any other equipment that could give rise to pressure fluctuations
* the number of gauges required to safely monitor facility pressures will depend on the number of rooms, the number of independent supply and exhaust systems, and layout
* the audible alarm is to indicate loss of negative room pressure in excess of two minutes
* where practicable other air-conditioning control switches and exhaust fan speed set point control should also be located adjacent to the gauges
* the HEPA filter gauge can also be mounted with the room gauges
* if there is a closable door between two rooms they are separate.

Section 10: Animals and animal containment facilities (only 10.8.1 (b) associated notes only and 10.9.2 (c) – paragraph 5 sentences 2 and 3 and associated note).

* 1. In substitution for 10.9.2 (c) paragraph 1, doors to the facility airlock must be fitted with automatic closers unless ‘door open’ alarms are fitted.

### Plant laboratory containment—level 3 (BC3) facilities

The BIP of plant laboratory containment approval must also meet Part 3.1 about the BC3 requirements and the requirements for approval, the General requirements and the Specific requirements for Plant Containment.

* 1. The following standards from AS/NZS 2243.3:2002 are the minimum for work with plant goods at the BC3 level.

Section 4.8: PC2 requirements (Only 4.8.3 (b)).

* 1. In addition to 4.8.3 (b) the following apply:

Large HVAC heat exchangers (e.g. chilled beams and some other natural or forced convectors) located within containment occupancies may be vulnerable to contamination, if they are not protected by efficient air filters. The likely contamination and implications for containment should be carefully assessed where such devices are proposed or are used.

Section 4.9: PC3 requirements (only 4.9.2 (a) – excluding the requirements for doors to open outwards, and be self-closing. The requirement for doors to open outwards will be assessed as flexible depending on the context of the facility and rooms within. 4.9.2 (b), (c) and 4.9.3 (a), (b), (c), (d), (e), (f) and (i)).

* 1. In substitution for 4.9.3 (g) the following clause will be applied. Each separate room within a facility must have:
* a room pressure gauge that can be viewed by personnel prior to entering the facility
* an audible alarm.
  1. Notes:
* the room pressure gauge must measure the differential pressure relative to the adjacent rooms external to the facility. The reference pressure should be taken from adjacent rooms, normally the access space not the ceiling space. The adjacent space should not be subject to significant variations in ambient static pressure due to external influences such as ambient winds, ventilation equipment, fans, or any other equipment that could give rise to pressure fluctuations
* the number of gauges required to safely monitor facility pressures will depend on the number of rooms, the number of independent supply and exhaust systems, and layout
* the audible alarm is to indicate loss of negative room pressure in excess of two minutes
* where practicable other air-conditioning control switches and exhaust fan speed set point control should also be located adjacent to the gauges
* where HEPA filter gauge can also be mounted with the room gauges
* if there is a closable door between two rooms they are separate.

## Additional requirements

### Specific requirements for BC3 microbiological, animal and plant facilities

* 1. A facility must be constructed so that each room within achieves upon commissioning an air leakage rate, at a differential pressure of 200 Pa, of no more than 120L/min. An air leakage test must be provided for new or refurbished facilities.

Note: A small ante-room/airlock of less than 10 m2 floor area may be added with the adjacent room for the purpose of this requirement.

* 1. Where drinking fountains are provided they must be of hands-free operation and be within a designated area where goods are not handled, stored or treated.
  2. Write-up areas may be approved as part of a BC3 facility. To be eligible for approval, these areas must comply with BC3 requirements, be constructed such that horizontal surfaces are minimised (for example, minimal shelving), must not be used for generic office functions and should hold only essential reference materials (such as technical equipment manuals).
  3. HEPA filters must be mounted in gas tight housings located as close as possible to the containment facility to minimize the length of potentially contaminated ductwork. The interconnecting ductwork between the containment room and the HEPA filter housing must also be of gastight (stainless steel) construction.
  4. The design of the filter housing must facilitate the testing of the integrity of the HEPA filter element and mounting, and the periodic gaseous decontamination of the filter element and housing.

Note: Housings should be placed in fully accessible locations with clear access to facilitate filter integrity testing, physical handling of filter elements and operation of isolating valves. Installations in false ceilings should be avoided.

* 1. To enable testing and gaseous decontamination filter housings should incorporate the following features:
* sealed access doors for filter maintenance and integrity testing
* gastight isolating valves on the air inlet and outlet ducts
* secure filter element clamping and mounting tracks
* upstream and downstream valved ports
* upstream and downstream valved pressure tappings to permit monitoring of the filter air flow pressure drop
* a differential pressure gauge incorporating a magnetically coupled indicating mechanism and a sealed differential pressure diaphragm
* a facility to introduce a test airflow and cold generated aerosol to establish the integrity of the filter element and its mounting.
  1. Waste piping must be installed such that the length of horizontal piping is minimised. The pipe path should have the maximum practical fall (preferably vertical). The pipe should be routed via plant rooms and accessible building risers. The route should avoid ceiling spaces and occupied areas unless this is impractical.
  2. Piping must be conservatively selected to suit the fluid flow and pressure applicable. The pipe material should resist degradation from exposure to waste products or likely cleaning and disinfection agents

Note*:* Fully welded 316l stainless steel piping of 1.0 mm thickness is recommended. Single skin piping is considered satisfactory.

* 1. Piping must be physically protected where exposed to mechanical damage.
  2. Piping must be labelled ‘Biosecurity Containment Pipework – Do not disturb’ throughout its length. Biohazard warning symbols should also be provided at regular intervals.
  3. Piping should be capable of being visually inspected throughout its length. Double skin pipe construction is recommended in any locations where visual inspection is unable to be undertaken. However this contingency should be avoided, where practical.
  4. A filtration system is required for removal of solids prior to liquid waste entering the holding tank. The filtration system must be capable of being removed for cleaning and decontamination. Solids removed from waste water are deemed to be biosecurity waste and must be treated as such. Steam sterilisation is recommended.

Note*:* The efficacy of the waste treatment system in sterilising particulate matter should be considered when determining the filtration system screen size to be utilised for solids removal.

* 1. The floor of the waste treatment plant room must be fully bunded to a volume that ensures retention of all waste in the event of a holding tank failure at full capacity. The bunded space should drain to a sealed sump. A submersible sump pump must be provided, along with flexible hosing to discharge the spillage after chemical disinfection

Note: Consideration should be given to the provision of a safety shower and eyewash station within the plant room.

* 1. A continuous flow alarm is desirable for water services connected to appliances draining to the treatment system. A suitable timer can warn users where water flow to a single appliance exceeds a reasonable period of time.
  2. The vents to the waste system must be fitted with sterile filters.

Note: Filters must be carefully selected to ensure that sufficient air is passed to ensure traps are not compromised. 0.2 micron hydrophobic membrane filters are recommended. These must be capable of being decontaminated, preferably by steam sterilisation.

* 1. Where toilets are installed, it is recommended that these are urine only systems. Where urine only toilets are installed a ‘Lady-San’ or similar disposal station for small quantities of toilet paper and other minor solids should be supplied. Solid waste collected from disposal stations must be treated as biosecurity waste.

### Specific requirements for BC3 animal and plant facilities

* 1. Any openings in the walls, ceiling or roof, such as vents, drainage outlets and air conditioning or ventilation inlets and outlets, must be screened at the containment boundary with fine mesh screens having an aperture size small enough to prevent entry or egress of insects.
  2. An aperture size small enough to prevent entry or egress of insects will require a maximum aperture size of 0.25 mm or 250 microns (um). Suitable material includes stainless steel mesh of 0.16 mm wire gauge (0.25 mm aperture).
  3. Screens must be of suitable material to withstand the air flow load, to remain undamaged following cleaning and be resistant to attack by insects or corrosion.

The size of insects to which plants and animals subject to biosecurity control held within the facility are potential hosts must also be considered when determining the appropriate screen size.

* 1. The facility should have an automatically-starting emergency power source to ensure continuing operation of the ventilation system, room access and if applicable shower controls.
  2. Sufficient uninterruptible power should be provided for essential equipment such as biological safety cabinets.
  3. The emergency power source would need to be designed to enable an automatic changeover as soon as practicable. Any such system would need to ensure emergency lighting and communication systems are reactivated within the shortest practical timeframe.

An uninterruptible power supply for biosafety cabinets would ensure seamless supply is maintained.